Intended Use:

The Medisafe Sonic Irrigator Semi-Automatic is principally intended to clean hollow (cannulated) instruments, surgical instruments, and associated products. This is achieved through a combination of ultrasonic activity and low-frequency Cannulated Pulse Enhancement (CPE) fluid re-circulation to dislodge and flush away unwanted debris from both the outside and the inside (hollow instruments only) of items placed in the wash tank.

Product features:

Technical Specifications

| Equipment designation | Reliance S Ultrasonic Cleaner |
| --- | --- |
| Classifications | Medical device class I Class 1 electrical shock protection, externally powered No applied parts IP rating: IPX0 Not suitable for use in oxygen-rich environment Designed for continuous use |
| Intended purpose | Cleaning of surgical instruments, and associated products. |
| Operating process | Ultrasonic activity and water heating. |
| Ultrasonic system | MED511030, MED511020 (10 Litres): 4 piezo-electric transducers MED512030, MED512020 (20 Litres): 6 piezo-electric transducers Typically operating at 38 to 40 kHz dependent on load. 40 W power per transducer. Pulsed degas, auto-tune, frequency sweep. |
| Water temperature monitoring | Measurement range: +2 ºC to 60 ºC (+36 ºF to 140 ºF) Temp. display: +5 ºC to 60 ºC (+40 ºF to 140 ºF), accuracy ±0.5. |
| Drainage system | Manually operated valve and drain pump |
| Drain outlet | 3/4 in. hose tail |
| Effluent temperature | Max. +60 ºC (+140 ºF) |
| Maximum drain flow rate | less than 15 litre/min. (4 gal/min.) |
| User interface | 4.3 in. TFT touch-screen display Time display: minutes and seconds |
| Audio | 1 W speaker |
| Internal clock accuracy | ±0.5 sec |
| Internal data storage | Micro SD flash card, 4 GB (storage of >10,000 cycles) |
| Real-time clock battery | CR1220 (5012LC), 35 to 40 mAh, 0.1 mA |
| Altitude range | Interior only, altitude up to 2000 m (6562 ft) |
| Ambient operating temperature range | +5 ºC to 40 ºC (+41 ºF to 104 ºF) |
| Maximum humidity | 80% relative humidity for up to +31 ºC (+88 ºF) decreasing linearly to 50% relative humidity at +40 ºC (+104 ºF) |
| Noise generated during operation | 71 dBA at distance of 1 m (3.3 ft) from front of machine |
| Heat output during operation | 205 W (700 BTU/hr) during wash cycle |

|  | MED511020 & MED511030 | MED512020 & MED512030 |
| --- | --- | --- |
| External height | 319 mm (12.6 in.) | 319 mm (12.6 in.) |
| External width | 357 mm (14.1 in.) | 553 mm (21.8 in.) |
| External depth | 380 mm (15.0 in.) | 440 mm (17.3 in.) |
| Net weight (dry) | 8.8 kg (19.4 lb) | 16.5 kg (36.2 lb) |
| Operating weight (tank filled) | 18.4 kg (40.5 lb) | 32.4 kg (71.2 lb) |
|  | MED511020 & MED511030 | MED512020 & MED512030 |
| Wash tank length | 290 mm (11.4 in.) | 495 mm (19.4 in.) |
| Wash tank width | 230 mm (9.0 in.) | 282 mm (11.1 in.) |
| Wash tank depth | 200 mm (7.8 in.) | 150 mm (5.9 in.) |
| Wash tank liquid volume | 10 L (2.5 gal) | 20 L (5.2 gal) |
|  | MED511020 & MED511030 | MED512020 & MED512030 |
| Basket length | 228 mm (8.9 in.) | 431 mm (16.9 in.) |
| Basket width | 187 mm (7.3 in.) | 235 mm (9.2 in.) |
| Basket depth | 154 mm (6.0 in.) | 121 mm (4.7 in.) |
| Maximum load weight | 3.0 kg (6.6 lb) | 4.0 kg (8.8 lb) |

Daily Maintenance

|  | Before every cycle Perform the following procedure after each wash cycle: Replace drain filter in wash tank Carefully remove the small, caged filter located in the drain inside the wash tank. Replace with a new item. |
| --- | --- |
|  | Wipe water level sensor in tank Carefully wipe any excess moisture or dirt from the water level sensor using a lint free cloth, soft brush, or an alcohol impregnated swab. Drain wash tank in the following way: Connect hose to drain tap and place other end into a suitable drain Open tap on drain valve Press and hold the ‘speed drain’ button on front, right-hand side of the unit until wash tank is empty |

Daily Checks - Remove Lid and make sure Vapor Barrier and Lid Magnets are clean and undamaged (Image 8.1). VAPOR BARRIER LID MAGNET LID MAGNET Image 8.1: Location of Lid Magnets and Vapor Barrier on Lid - Remove Basket and make sure it is clean and undamaged. - Clean Water Level Sensor with sterile wipe or soft nylon brush (Image 8.2). WATER LEVEL SENSOR Image 8.2: Location of Water Level Sensor and Drain Filter in Tank - Inspect Drain Filter and replace if clogged (Image 8.2). - With clean cloth and disinfectant, wipe following; - All interior surfaces of Tank. - All external surfaces of unit. - All surfaces of Lid. - With clean wet cloth, remove any trace of disinfectant from internal and external surfaces of unit and Lid.

Additional Information: